## Gain DIP Switch Configuration
FUTEK Advanced Sensor Technology, Inc.
IAA200 Gain and Shunt Settings
1. Excitation | 10 | VDC | Directions: 1. Select excitation voltage to sensor from amplifier. 5V or 10V can be set on the amplifier board. 2. Enter the sensitivity (mV/V) of the sensor as found on sensor calibration certificate. 3. Choose amplifier full output level. 4. The yellow rows in the table below reflect gain DIP switch combinations that refelect a gain multiple to amplify the sensors output to the select amplifier full output.
2. Output | 2 | mV/V
3. Amplifier Full Output | 20 | mA
Switch Down | 0 | Off
Switch Up | 1 | On
Select the highlighted row that will provide the gain that results in the output voltage closest to the specified value.
Iout = (Excitation * Sensor mV/V * Gain) * (16/10) + 4
Below are DIP switch configurations that will satisfy the above input information.
SW3 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Sensitivity (mV/V) | 0.5 | 1 | 1.5 | 2 | 2.5 | 3 | 4 | 10
Resistance (Ohms) | 23.7 | 47.5 | 71.5 | 95.3 | 124 | 143 | 191 | 475 | Combined Resistance | Gain | Output (mA)
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 1 | 4.032
1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 475 | 106.2632 | 7.4004
2 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 191 | 262.7801 | 12.409
3 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 143 | 350.6503 | 15.2208
4 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 136.2237 | 368.0433 | 15.7774
5 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 124 | 404.2258 | 16.9352
6 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 109.911 | 455.9135 | 18.5892
7 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 98.3306 | 509.489 | 20.3036
8 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 95.3 | 525.659 | 20.8211
9 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 81.7754 | 612.4305 | 23.5978
10 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 79.3749 | 630.9221 | 24.1895
11 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 75.1873 | 666.0059 | 25.3122
12 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 71.5 | 700.3007 | 26.4096
13 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 69.7648 | 717.6936 | 26.9662
14 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 66.412 | 753.8762 | 28.124
15 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 64.9124 | 771.2691 | 28.6806
16 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 63.5777 | 787.4391 | 29.1981
17 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 62.1455 | 805.5639 | 29.778
18 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 58.2656 | 859.1393 | 31.4925
19 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 57.188 | 875.3093 | 32.0099
20 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 56.0725 | 892.7022 | 32.5665
21 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 53.886 | 928.8848 | 33.7243
22 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 52.0248 | 962.0808 | 34.7866
23 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 51.0427 | 980.5725 | 35.3783
24 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 49.2778 | 1015.6563 | 36.501
25 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 48.3958 | 1034.1479 | 37.0927
26 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 47.6667 | 1049.951 | 37.5984
27 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 47.5 | 1053.6316 | 37.7162
28 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 46.8892 | 1067.344 | 38.155
29 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 45.3504 | 1103.5265 | 39.3128
30 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 44.6461 | 1120.9194 | 39.8694
31 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 44.0106 | 1137.0894 | 40.3869
32 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 43.3195 | 1155.2142 | 40.9669
33 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 43.1818 | 1158.8947 | 41.0846
34 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 42.0286 | 1190.6649 | 42.1013
35 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 41.3979 | 1208.7897 | 42.6813
36 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 40.851 | 1224.9597 | 43.1987
37 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 40.2786 | 1242.3526 | 43.7553
38 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 39.1379 | 1278.5351 | 44.9131
39 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 38.6122 | 1295.928 | 45.4697
40 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 38.1466 | 1311.7312 | 45.9754
41 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 38.0398 | 1315.4117 | 46.0932
42 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 37.616 | 1330.2228 | 46.5671
43 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 36.6487 | 1365.3066 | 47.6898
44 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 36.1586 | 1383.7983 | 48.2815
45 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 35.6562 | 1403.2819 | 48.905
46 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 35.3109 | 1416.9943 | 49.3438
47 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 35.2193 | 1420.6748 | 49.4616
48 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 34.4311 | 1453.1769 | 50.5017
49 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 34.344 | 1456.8574 | 50.6194
50 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 34.0236 | 1470.5698 | 51.0582
51 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 33.6533 | 1486.7398 | 51.5757
52 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 33.1665 | 1508.5451 | 52.2734
53 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 32.482 | 1540.3152 | 53.2901
54 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 32.104 | 1558.44 | 53.8701
55 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 32.0283 | 1562.1205 | 53.9879
56 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 31.7741 | 1574.61 | 54.3875
57 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 31.6999 | 1578.2906 | 54.5053
58 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 31.4267 | 1592.0029 | 54.9441
59 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 30.7279 | 1628.1855 | 56.1019
60 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 30.4029 | 1645.5784 | 56.6585
61 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 30.047 | 1665.062 | 57.282
62 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 29.7819 | 1679.8732 | 57.7559
63 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 29.7167 | 1683.5537 | 57.8737
64 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 29.1723 | 1714.957 | 58.8786
65 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 29.1098 | 1718.6375 | 58.9964
66 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 28.8609 | 1733.4486 | 59.4704
67 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 28.5399 | 1752.9323 | 60.0938
68 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 28.2594 | 1770.3252 | 60.6504
69 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 27.693 | 1806.5077 | 61.8082
70 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 27.4843 | 1820.2201 | 62.247
71 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 27.4288 | 1823.9006 | 62.3648
72 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 27.2422 | 1836.3901 | 62.7645
73 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 27.1876 | 1840.0707 | 62.8823
74 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 26.9223 | 1858.1954 | 63.4623
75 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 26.4695 | 1889.9656 | 64.4789
76 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 26.1674 | 1911.7709 | 65.1767
77 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 25.9479 | 1927.9409 | 65.6941
78 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 25.7645 | 1941.6533 | 66.1329
79 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 25.7157 | 1945.3338 | 66.2507
80 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 25.2929 | 1977.8358 | 67.2907
81 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 25.2459 | 1981.5164 | 67.4085
82 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 25.0723 | 1995.2287 | 67.8473
83 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 24.8298 | 2014.7124 | 68.4708
84 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 24.6038 | 2033.2041 | 69.0625
85 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 24.1863 | 2068.2878 | 70.1852
86 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 24.0142 | 2083.099 | 70.6592
87 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 23.9719 | 2086.7795 | 70.7769
88 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 23.7916 | 2102.5826 | 71.2826
89 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 23.7 | 2110.7046 | 71.5425
90 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 23.5963 | 2119.9755 | 71.8392
91 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 23.2002 | 2156.1581 | 72.9971
92 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 23.0144 | 2173.551 | 73.5536
93 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 22.8444 | 2189.721 | 74.0711
94 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 22.6568 | 2207.8458 | 74.6511
95 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 22.5737 | 2215.9678 | 74.911
96 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 22.3352 | 2239.6159 | 75.6677
97 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 22.2986 | 2243.2965 | 75.7855
98 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 22.1198 | 2261.4212 | 76.3655
99 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 21.9627 | 2277.5912 | 76.8829
100 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 21.7961 | 2294.9842 | 77.4395
101 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 21.4577 | 2331.1667 | 78.5973
102 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 21.3322 | 2344.8791 | 79.0361
103 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 21.2987 | 2348.5596 | 79.1539
104 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 21.1563 | 2364.3627 | 79.6596
105 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 21.0838 | 2372.4847 | 79.9195
106 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 20.992 | 2382.8544 | 80.2513
107 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 20.6873 | 2417.9382 | 81.374
108 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 20.5303 | 2436.4299 | 81.9658
109 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 20.3305 | 2460.355 | 82.7314
110 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 20.2542 | 2469.6259 | 83.028
111 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 20.1878 | 2477.7479 | 83.2879
112 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 19.9616 | 2505.8084 | 84.1859
113 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 19.8971 | 2513.9304 | 84.4458
114 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 19.824 | 2523.2013 | 84.7424
115 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 19.6977 | 2539.3714 | 85.2599
116 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 19.4961 | 2565.6181 | 86.0998
117 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 19.2905 | 2592.9468 | 86.9743
118 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 19.1566 | 2611.0716 | 87.5543
119 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 19.0971 | 2619.1936 | 87.8142
120 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 19.0386 | 2627.2416 | 88.0717
121 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 18.9799 | 2635.3636 | 88.3316
122 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 18.9134 | 2644.6345 | 88.6283
123 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 18.658 | 2680.8171 | 89.7861
124 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 18.5377 | 2698.21 | 90.3427
125 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 18.3747 | 2722.1351 | 91.1083
126 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 18.3049 | 2732.5048 | 91.4402
127 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 18.2507 | 2740.6268 | 91.7001
128 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 18.0728 | 2767.5885 | 92.5628
129 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 18.0199 | 2775.7106 | 92.8227
130 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 17.9528 | 2786.0802 | 93.1546
131 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 17.7999 | 2810.0053 | 93.9202
132 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 17.6904 | 2827.3983 | 94.4767
133 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 17.4668 | 2863.5808 | 95.6346
134 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 17.4104 | 2872.8517 | 95.9313
135 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 17.3613 | 2880.9737 | 96.1912
136 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 17.3129 | 2889.0217 | 96.4487
137 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 17.2643 | 2897.1437 | 96.7086
138 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 17.157 | 2915.2685 | 97.2886
139 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 16.9976 | 2942.5972 | 98.1631
140 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 16.8472 | 2968.844 | 99.003
141 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 16.756 | 2985.014 | 99.5204
142 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 16.7041 | 2994.2849 | 99.8171
143 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 16.6589 | 3002.4069 | 100.077
144 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 16.5046 | 3030.4674 | 100.975
145 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 16.4604 | 3038.5894 | 101.2349
146 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 16.4103 | 3047.8603 | 101.5315
147 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 16.2825 | 3071.7854 | 102.2971
148 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 16.185 | 3090.2771 | 102.8889
149 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 16.0033 | 3125.3609 | 104.0115
150 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 15.9503 | 3135.7306 | 104.3434
151 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 15.9091 | 3143.8526 | 104.6033
152 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 15.8295 | 3159.6557 | 105.109
153 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 15.8111 | 3163.3362 | 105.2268
154 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 15.7428 | 3177.0486 | 105.6656
155 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 15.5655 | 3213.2311 | 106.8234
156 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 15.4817 | 3230.6241 | 107.38
157 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 15.4046 | 3246.7941 | 107.8974
158 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 15.319 | 3264.9188 | 108.4774
159 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 15.3018 | 3268.5994 | 108.5952
160 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 15.1918 | 3292.2475 | 109.3519
161 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 15.1544 | 3300.3695 | 109.6118
162 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 15.0716 | 3318.4943 | 110.1918
163 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 14.9985 | 3334.6643 | 110.7093
164 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 14.9207 | 3352.0572 | 111.2658
165 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 14.7613 | 3388.2398 | 112.4237
166 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 14.721 | 3397.5107 | 112.7203
167 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 14.6859 | 3405.6327 | 112.9802
168 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 14.618 | 3421.4358 | 113.4859
169 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 14.6023 | 3425.1163 | 113.6037
170 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 14.5394 | 3439.9275 | 114.0777
171 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 14.3926 | 3475.0113 | 115.2004
172 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 14.3164 | 3493.5029 | 115.7921
173 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 14.237 | 3512.9866 | 116.4156
174 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 14.1816 | 3526.699 | 116.8544
175 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 14.1668 | 3530.3795 | 116.9721
176 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 14.0375 | 3562.8815 | 118.0122
177 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 14.023 | 3566.562 | 118.13
178 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 13.9693 | 3580.2744 | 118.5688
179 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 13.9065 | 3596.4444 | 119.0862
180 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 13.8227 | 3618.2497 | 119.784
181 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 13.7023 | 3650.0199 | 120.8006
182 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 13.6346 | 3668.1447 | 121.3806
183 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 13.6209 | 3671.8252 | 121.4984
184 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 13.5747 | 3684.3147 | 121.8981
185 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 13.5612 | 3687.9952 | 122.0158
186 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 13.5109 | 3701.7076 | 122.4546
187 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 13.3801 | 3737.8901 | 123.6125
188 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 13.3181 | 3755.283 | 124.1691
189 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 13.2494 | 3774.7667 | 124.7925
190 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 13.1976 | 3789.5778 | 125.2665
191 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 13.1848 | 3793.2583 | 125.3843
192 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 13.0765 | 3824.6616 | 126.3892
193 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 13.0639 | 3828.3421 | 126.5069
194 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 13.0135 | 3843.1533 | 126.9809
195 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 12.9479 | 3862.6369 | 127.6044
196 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 12.8898 | 3880.0298 | 128.161
197 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 12.7707 | 3916.2124 | 129.3188
198 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 12.7261 | 3929.9248 | 129.7576
199 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 12.7142 | 3933.6053 | 129.8754
200 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 12.674 | 3946.0948 | 130.275
201 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 12.6622 | 3949.7753 | 130.3928
202 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 12.6043 | 3967.9001 | 130.9728
203 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 12.5042 | 3999.6702 | 131.9894
204 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 12.4363 | 4021.4755 | 132.6872
205 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 12.3865 | 4037.6455 | 133.2047
206 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 12.3446 | 4051.3579 | 133.6435
207 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 12.3334 | 4055.0385 | 133.7612
208 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 12.2353 | 4087.5405 | 134.8013
209 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 12.2243 | 4091.221 | 134.9191
210 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 12.1834 | 4104.9334 | 135.3579
211 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 12.1259 | 4124.417 | 135.9813
212 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 12.0717 | 4142.9087 | 136.5731
213 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 11.9703 | 4177.9925 | 137.6958
214 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 11.928 | 4192.8036 | 138.1697
215 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 11.9176 | 4196.4842 | 138.2875
216 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 11.8729 | 4212.2873 | 138.7932
217 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 11.824 | 4229.6802 | 139.3498
218 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 11.7237 | 4265.8627 | 140.5076
219 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 11.6761 | 4283.2556 | 141.0642
220 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 11.6322 | 4299.4256 | 141.5816
221 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 11.5833 | 4317.5504 | 142.1616
222 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 11.4987 | 4349.3206 | 143.1783
223 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 11.489 | 4353.0011 | 143.296
224 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 11.4413 | 4371.1259 | 143.876
225 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 11.3991 | 4387.2959 | 144.3935
226 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 11.3541 | 4404.6888 | 144.95
227 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 11.2616 | 4440.8713 | 146.1079
228 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 11.2269 | 4454.5837 | 146.5467
229 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 11.2176 | 4458.2643 | 146.6645
230 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 11.178 | 4474.0674 | 147.1702
231 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 11.132 | 4492.559 | 147.7619
232 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 11.0457 | 4527.6428 | 148.8846
233 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 11.0008 | 4546.1345 | 149.4763
234 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 10.921 | 4579.3305 | 150.5386
235 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 10.8354 | 4615.5131 | 151.6964
236 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 10.7947 | 4632.906 | 152.253
237 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 10.7571 | 4649.076 | 152.7704
238 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 10.6346 | 4702.6515 | 154.4848
239 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 10.5937 | 4720.7762 | 155.0648
240 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 10.5576 | 4736.9462 | 155.5823
241 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 10.5189 | 4754.3392 | 156.1389
242 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 10.4395 | 4790.5217 | 157.2967
243 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 10.4017 | 4807.9146 | 157.8533
244 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 10.328 | 4842.2094 | 158.9507
245 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 10.2537 | 4877.2932 | 160.0734
246 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 10.215 | 4895.7849 | 160.6651
247 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 10.037 | 4982.5563 | 163.4418
248 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 10.0045 | 4998.7263 | 163.9592
249 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 9.8984 | 5052.3018 | 165.6737
250 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 9.7982 | 5103.9895 | 167.3277
251 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 9.7292 | 5140.172 | 168.4855
252 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 9.6964 | 5157.565 | 169.0421
253 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 9.5339 | 5245.4352 | 171.8539
254 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 9.2576 | 5401.9522 | 176.8625
255 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 9.0806 | 5507.2153 | 180.2309

## Shunt DIP Switch Configuration
FUTEK Advanced Sensor Technology, Inc.
IAA200 Gain and Shunt Settings
1. Excitation | 10 | VDC | Directions: 1. Select excitation voltage to sensor from amplifier. 5V or 10V can be set on the amplifier board. 2. Enter the bridge resistance (Ω) of the sensor between the +/- Signal connections. 3. Enter the gain of the system. The Gain DIP Switch Configuration worksheet in Column K shows the gain level for different gain DIP switch settings. 4.Choose amplifier full output level.
2. Bridge Resistance | 350 | Ω
3. Gain | 509
4. Amplifier Full Output | 20 | mA
Switch Down | 0 | Off
Switch Up | 1 | On
Select the highlighted row that will provide the shunt value that results in the output voltage closest to the specified value.
Below are DIP switch configurations that will satisfy the above input information.
SW5 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Resistance (Ohms) | 30000 | 43700 | 60400 | 87600 | 100000 | 150000 | 300000 | 432000 | Combined Resistance | RT | V1-V2 | Output (mA)
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 432000 | 349.7167 | 0.0020246 | 5.6489
2 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 300000 | 349.5921 | 0.002915 | 6.3739
3 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 177049.1803 | 349.3095 | 0.0049372 | 8.0209
4 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 150000 | 349.1852 | 0.0058265 | 8.7451
5 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 111340.2062 | 348.9032 | 0.0078465 | 10.3902
6 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 100000 | 348.7793 | 0.0087347 | 11.1136
7 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 100000 | 348.7793 | 0.0087347 | 11.1136
8 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 87600 | 348.6072 | 0.0099687 | 12.1185
9 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 81203.0075 | 348.4979 | 0.010752 | 12.7567
10 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 81203.0075 | 348.4979 | 0.010752 | 12.7567
11 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 75000 | 348.3743 | 0.01164 | 13.4792
12 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 72831.4088 | 348.3261 | 0.011985 | 13.7608
13 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 67801.8576 | 348.2025 | 0.012872 | 14.483
14 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 63905.3254 | 348.0935 | 0.013655 | 15.1204
15 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 60400 | 347.9835 | 0.014445 | 15.7639
16 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 60000 | 347.9702 | 0.014541 | 15.8421
17 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 58604.0288 | 347.9221 | 0.014886 | 16.1234
18 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 55303.0303 | 347.7989 | 0.015772 | 16.8447
19 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 52991.0642 | 347.7035 | 0.016458 | 17.4033
20 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 52682.9268 | 347.6901 | 0.016554 | 17.4814
21 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 50277.4695 | 347.5804 | 0.017343 | 18.1242
22 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 50000 | 347.567 | 0.017439 | 18.2023
23 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 49026.8018 | 347.5191 | 0.017784 | 18.4832
24 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 46695.0959 | 347.3961 | 0.018669 | 19.2037
25 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 46695.0959 | 347.3961 | 0.018669 | 19.2037
26 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 45036.0388 | 347.3009 | 0.019354 | 19.7616
27 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 44813.278 | 347.2876 | 0.01945 | 19.8397
28 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 43700 | 347.2191 | 0.019943 | 20.2416
29 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 43060.8365 | 347.1781 | 0.020238 | 20.4817
30 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 42140.1465 | 347.117 | 0.020678 | 20.8403
31 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 42140.1465 | 347.117 | 0.020678 | 20.8403
32 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 40405.9041 | 346.9943 | 0.021562 | 21.56
33 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 39685.5161 | 346.9402 | 0.021952 | 21.8773
34 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 39157.6824 | 346.8993 | 0.022246 | 22.1172
35 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 38143.73 | 346.8177 | 0.022835 | 22.5966
36 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 37655.8603 | 346.7768 | 0.023129 | 22.8365
37 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 37655.8603 | 346.7768 | 0.023129 | 22.8365
38 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 36949.8992 | 346.7158 | 0.023569 | 23.1947
39 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 35750.2703 | 346.6067 | 0.024356 | 23.8356
40 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 35609.7561 | 346.5934 | 0.024452 | 23.9135
41 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 35049.0506 | 346.5395 | 0.024841 | 24.2305
42 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 34636.7054 | 346.4987 | 0.025135 | 24.4701
43 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 34636.7054 | 346.4987 | 0.025135 | 24.4701
44 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 33840.9912 | 346.4172 | 0.025723 | 24.949
45 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 33456.4254 | 346.3764 | 0.026017 | 25.1885
46 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 33017.8682 | 346.3288 | 0.026361 | 25.4685
47 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 32897.9762 | 346.3156 | 0.026457 | 25.5463
48 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 31943.6261 | 346.2067 | 0.027243 | 26.1865
49 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 31831.3953 | 346.1935 | 0.027338 | 26.2643
50 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 31382.6144 | 346.1396 | 0.027727 | 26.5809
51 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 31051.6194 | 346.0989 | 0.028021 | 26.8203
52 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 30410.5776 | 346.0176 | 0.028608 | 27.2986
53 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 30410.5776 | 346.0176 | 0.028608 | 27.2986
54 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 30099.6678 | 345.977 | 0.028902 | 27.5378
55 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 30000 | 345.9638 | 0.028998 | 27.6156
56 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 29744.2312 | 345.9295 | 0.029245 | 27.8175
57 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 29646.8996 | 345.9162 | 0.029341 | 27.8952
58 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 29155.5217 | 345.8482 | 0.029832 | 28.2955
59 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 28869.6244 | 345.8076 | 0.030126 | 28.5347
60 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 28410.6164 | 345.7407 | 0.03061 | 28.9286
61 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 28410.6164 | 345.7407 | 0.03061 | 28.9286
62 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 28139.0734 | 345.7001 | 0.030903 | 29.1677
63 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 28051.9481 | 345.6869 | 0.030999 | 29.2454
64 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 27611.6259 | 345.619 | 0.03149 | 29.6454
65 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 27355.0725 | 345.5784 | 0.031783 | 29.8844
66 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 27312.2294 | 345.5716 | 0.031833 | 29.9248
67 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 27272.7273 | 345.5652 | 0.031879 | 29.9621
68 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 27061.1841 | 345.531 | 0.032126 | 30.1637
69 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 26573.0208 | 345.45 | 0.032713 | 30.6412
70 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 26335.3216 | 345.4095 | 0.033006 | 30.8801
71 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 26335.3216 | 345.4095 | 0.033006 | 30.8801
72 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 25952.83 | 345.3427 | 0.033489 | 31.2736
73 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 25726.0494 | 345.3022 | 0.033782 | 31.5124
74 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 25653.2067 | 345.289 | 0.033878 | 31.59
75 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 25355.2354 | 345.2344 | 0.034273 | 31.912
76 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 25284.4744 | 345.2213 | 0.034368 | 31.9896
77 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 25033.1887 | 345.174 | 0.034711 | 32.2686
78 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 25000 | 345.1677 | 0.034757 | 32.3059
79 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 24822.13 | 345.1335 | 0.035004 | 32.5073
80 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 24822.13 | 345.1335 | 0.035004 | 32.5073
81 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 24410.7924 | 345.0527 | 0.03559 | 32.9842
82 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 24210.0562 | 345.0122 | 0.035883 | 33.2228
83 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 23949.571 | 344.9588 | 0.03627 | 33.5383
84 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 23886.4286 | 344.9456 | 0.036365 | 33.6159
85 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 23632.3851 | 344.8921 | 0.036753 | 33.9319
86 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 23379.2783 | 344.8376 | 0.037148 | 34.2535
87 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 23319.1035 | 344.8245 | 0.037243 | 34.331
88 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 23105.1993 | 344.7773 | 0.037586 | 34.6097
89 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 23076.9231 | 344.771 | 0.037631 | 34.6469
90 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 23076.9231 | 344.771 | 0.037631 | 34.6469
91 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 22925.2822 | 344.7369 | 0.037878 | 34.8481
92 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 22573.9646 | 344.6562 | 0.038463 | 35.3245
93 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 22573.9646 | 344.6562 | 0.038463 | 35.3245
94 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 22402.1949 | 344.6159 | 0.038756 | 35.5628
95 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 22346.9388 | 344.6028 | 0.038851 | 35.6403
96 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 22178.9807 | 344.5626 | 0.039143 | 35.878
97 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 22124.8189 | 344.5495 | 0.039238 | 35.9554
98 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 21906.6937 | 344.496 | 0.039626 | 36.2711
99 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 21906.6937 | 344.496 | 0.039626 | 36.2711
100 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 21689.032 | 344.4417 | 0.04002 | 36.5923
101 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 21452.9504 | 344.3815 | 0.040457 | 36.9481
102 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 21452.9504 | 344.3815 | 0.040457 | 36.9481
103 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 21428.5714 | 344.3752 | 0.040503 | 36.9853
104 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 21297.7585 | 344.3412 | 0.040749 | 37.1862
105 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 21247.8103 | 344.3281 | 0.040844 | 37.2636
106 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 20994.2218 | 344.2607 | 0.041334 | 37.6621
107 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 20845.5729 | 344.2205 | 0.041626 | 37.9001
108 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 20797.7208 | 344.2074 | 0.041721 | 37.9775
109 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 20652.1673 | 344.1673 | 0.042012 | 38.2149
110 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 20415.879 | 344.1009 | 0.042495 | 38.6076
111 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 20226.7064 | 344.0467 | 0.042889 | 38.9285
112 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 20226.7064 | 344.0467 | 0.042889 | 38.9285
113 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 20044.2478 | 343.9934 | 0.043276 | 39.2436
114 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 20021.2351 | 343.9866 | 0.043325 | 39.2838
115 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 20000 | 343.9803 | 0.043371 | 39.3209
116 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 19886.001 | 343.9464 | 0.043617 | 39.5217
117 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 19842.4483 | 343.9334 | 0.043712 | 39.5989
118 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 19663.7067 | 343.8792 | 0.044106 | 39.9197
119 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 19621.1213 | 343.8661 | 0.044201 | 39.997
120 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 19449.3783 | 343.8129 | 0.044587 | 40.312
121 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 19322.0281 | 343.7729 | 0.044879 | 40.5492
122 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 19322.0281 | 343.7729 | 0.044879 | 40.5492
123 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 19155.459 | 343.7197 | 0.045265 | 40.8641
124 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 19115.0442 | 343.7067 | 0.04536 | 40.9413
125 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 18949.1126 | 343.6526 | 0.045754 | 41.2619
126 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 18807.6243 | 343.6057 | 0.046095 | 41.5396
127 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 18788.8843 | 343.5994 | 0.04614 | 41.5767
128 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 18768.6625 | 343.5926 | 0.04619 | 41.6168
129 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 18750 | 343.5864 | 0.046235 | 41.6539
130 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 18611.4587 | 343.5395 | 0.046576 | 41.9316
131 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 18454.1188 | 343.4855 | 0.046969 | 42.2519
132 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 18416.6064 | 343.4724 | 0.047064 | 42.3291
133 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 18265.221 | 343.4194 | 0.047451 | 42.6438
134 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 18265.221 | 343.4194 | 0.047451 | 42.6438
135 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 18152.8612 | 343.3794 | 0.047742 | 42.8807
136 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 18005.7634 | 343.3263 | 0.048128 | 43.1953
137 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 17970.0499 | 343.3133 | 0.048223 | 43.2724
138 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 17823.3252 | 343.2594 | 0.048616 | 43.5926
139 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 17788.3311 | 343.2463 | 0.04871 | 43.6697
140 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 17698.0939 | 343.2126 | 0.048956 | 43.87
141 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 17681.4988 | 343.2063 | 0.049002 | 43.907
142 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 17663.5893 | 343.1996 | 0.049051 | 43.9471
143 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 17524.2837 | 343.1466 | 0.049437 | 44.2615
144 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 17524.2837 | 343.1466 | 0.049437 | 44.2615
145 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 17384.7198 | 343.0926 | 0.04983 | 44.5815
146 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 17216.9811 | 343.0267 | 0.050311 | 44.9729
147 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 17117.1126 | 342.9868 | 0.050601 | 45.2095
148 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 17084.8341 | 342.9738 | 0.050696 | 45.2866
149 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 16986.2614 | 342.9339 | 0.050987 | 45.5238
150 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 16823.8048 | 342.867 | 0.051474 | 45.9206
151 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 16792.622 | 342.8541 | 0.051569 | 45.9976
152 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 16712.1814 | 342.8204 | 0.051814 | 46.1977
153 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 16697.383 | 342.8141 | 0.05186 | 46.2347
154 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 16697.383 | 342.8141 | 0.05186 | 46.2347
155 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 16557.112 | 342.7545 | 0.052295 | 46.5888
156 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 16432.4733 | 342.7007 | 0.052687 | 46.9084
157 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 16432.4733 | 342.7007 | 0.052687 | 46.9084
158 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 16311.8433 | 342.6479 | 0.053073 | 47.2223
159 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 16282.5279 | 342.6349 | 0.053167 | 47.2993
160 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 16193.1778 | 342.5951 | 0.053457 | 47.5357
161 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 16164.2869 | 342.5822 | 0.053552 | 47.6127
162 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 16076.023 | 342.5423 | 0.053843 | 47.8495
163 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 16076.023 | 342.5423 | 0.053843 | 47.8495
164 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 15902.4745 | 342.4627 | 0.054424 | 48.3229
165 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 15830.3176 | 342.4291 | 0.054669 | 48.5227
166 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 15830.3176 | 342.4291 | 0.054669 | 48.5227
167 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 15817.0391 | 342.4229 | 0.054715 | 48.5597
168 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 15718.3363 | 342.3763 | 0.055055 | 48.8364
169 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 15691.1135 | 342.3634 | 0.055149 | 48.9133
170 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 15579.128 | 342.3097 | 0.055541 | 49.2326
171 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 15470.6601 | 342.257 | 0.055926 | 49.5462
172 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 15444.2877 | 342.244 | 0.05602 | 49.6231
173 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 15337.8679 | 342.1914 | 0.056405 | 49.9361
174 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 15258.3763 | 342.1517 | 0.056695 | 50.1727
175 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 15101.9468 | 342.0722 | 0.057276 | 50.6454
176 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 15101.9468 | 342.0722 | 0.057276 | 50.6454
177 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 15036.8569 | 342.0387 | 0.057521 | 50.845
178 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 15024.8756 | 342.0325 | 0.057566 | 50.882
179 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 1 | 14935.784 | 341.986 | 0.057906 | 51.1584
180 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 14911.2023 | 341.9731 | 0.058 | 51.2352
181 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 14810.0365 | 341.9195 | 0.058392 | 51.5541
182 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 14785.8666 | 341.9066 | 0.058486 | 51.6309
183 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 14711.9799 | 341.8669 | 0.058776 | 51.8673
184 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 14591.8421 | 341.8015 | 0.059254 | 52.2568
185 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 14591.8421 | 341.8015 | 0.059254 | 52.2568
186 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 14519.8771 | 341.7619 | 0.059545 | 52.4931
187 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 14378.1531 | 341.6826 | 0.060124 | 52.9653
188 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 14319.1406 | 341.6491 | 0.060369 | 53.1647
189 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 14308.2754 | 341.6429 | 0.060415 | 53.2016
190 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 14296.5453 | 341.6363 | 0.060463 | 53.2415
191 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 14227.4566 | 341.5966 | 0.060754 | 53.4777
192 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 14113.3078 | 341.5303 | 0.061239 | 53.8729
193 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 14091.3568 | 341.5174 | 0.061333 | 53.9497
194 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 14024.2325 | 341.4778 | 0.061623 | 54.1858
195 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 14024.2325 | 341.4778 | 0.061623 | 54.1858
196 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 13915.0227 | 341.4126 | 0.062101 | 54.5748
197 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 13849.5639 | 341.373 | 0.062391 | 54.8108
198 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 13741.3752 | 341.3067 | 0.062876 | 55.2059
199 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 13720.5651 | 341.2939 | 0.06297 | 55.2825
200 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 13666.8171 | 341.2605 | 0.063214 | 55.4817
201 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 13646.232 | 341.2476 | 0.063308 | 55.5584
202 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 13583.2719 | 341.2081 | 0.063598 | 55.7943
203 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 13583.2719 | 341.2081 | 0.063598 | 55.7943
204 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 13459.1639 | 341.1291 | 0.064177 | 56.2658
205 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 13397.9143 | 341.0896 | 0.064467 | 56.5016
206 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 13317.7543 | 341.0373 | 0.06485 | 56.8135
207 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 13298.2066 | 341.0245 | 0.064944 | 56.8902
208 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 13139.5248 | 340.9189 | 0.065718 | 57.5205
209 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 13120.4964 | 340.9061 | 0.065812 | 57.5971
210 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 13052.5069 | 340.8599 | 0.06615 | 57.8726
211 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 12994.8946 | 340.8205 | 0.066439 | 58.1083
212 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 12881.2606 | 340.7416 | 0.067018 | 58.5792
213 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 12881.2606 | 340.7416 | 0.067018 | 58.5792
214 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 12825.1469 | 340.7022 | 0.067307 | 58.8148
215 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 12751.6753 | 340.6501 | 0.067689 | 59.1263
216 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 12733.753 | 340.6372 | 0.067783 | 59.2029
217 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 12588.1823 | 340.5319 | 0.068557 | 59.8325
218 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 12508.2917 | 340.4731 | 0.068988 | 60.1842
219 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 12508.2917 | 340.4731 | 0.068988 | 60.1842
220 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 12455.3737 | 340.4337 | 0.069277 | 60.4195
221 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 12350.9416 | 340.355 | 0.069855 | 60.8899
222 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 12299.3439 | 340.3157 | 0.070144 | 61.1252
223 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 12231.7573 | 340.2637 | 0.070526 | 61.4364
224 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 12081.2459 | 340.1458 | 0.071392 | 62.1418
225 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 0 | 12081.2459 | 340.1458 | 0.071392 | 62.1418
226 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 12007.6414 | 340.0871 | 0.071824 | 62.4931
227 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 11958.8666 | 340.0478 | 0.072112 | 62.7282
228 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 11878.1146 | 339.9821 | 0.072595 | 63.1216
229 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 11862.5622 | 339.9693 | 0.072689 | 63.198
230 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 11752.5752 | 339.8782 | 0.073359 | 63.7439
231 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 1 | 11752.5752 | 339.8782 | 0.073359 | 63.7439
232 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 11613.5584 | 339.7606 | 0.074225 | 64.4484
233 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 11560.258 | 339.7147 | 0.074562 | 64.723
234 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 11545.5263 | 339.702 | 0.074655 | 64.7993
235 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 11425.7276 | 339.5972 | 0.075426 | 65.4272
236 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 11411.3366 | 339.5845 | 0.07552 | 65.5035
237 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 11309.5218 | 339.4936 | 0.07619 | 66.0487
238 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 11180.7315 | 339.3762 | 0.077054 | 66.7525
239 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 1 | 11131.3215 | 339.3305 | 0.07739 | 67.0267
240 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 11117.6621 | 339.3178 | 0.077484 | 67.1029
241 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 11006.5354 | 339.2133 | 0.078254 | 67.7301
242 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 10898.6597 | 339.1098 | 0.079016 | 68.3509
243 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 10779.0075 | 338.9927 | 0.079879 | 69.0538
244 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 10733.077 | 338.9471 | 0.080216 | 69.3277
245 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 10617.0136 | 338.8301 | 0.081078 | 70.0303
246 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 10617.0136 | 338.8301 | 0.081078 | 70.0303
247 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 10516.6035 | 338.7269 | 0.08184 | 70.6504
248 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 10362.3443 | 338.5646 | 0.083038 | 71.6261
249 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 10362.3443 | 338.5646 | 0.083038 | 71.6261
250 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 10254.1198 | 338.4479 | 0.0839 | 72.3279
251 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 10016.3674 | 338.1829 | 0.085857 | 73.9219
252 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 9915.2138 | 338.0665 | 0.086718 | 74.6229
253 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 9692.747 | 337.8021 | 0.088673 | 76.2151
254 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 9597.9933 | 337.686 | 0.089532 | 76.9152
255 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 9389.384 | 337.4222 | 0.091485 | 78.5056

## Sheet2 (hidden)
Excitation | IAA200 Gain and Shunt Settings v6 5/25/2021
5
10